## Calibration
DO SENSOR CALIBRATION
Fill the cream cells. Everything green calculates itself.
Date: |  | By:
STEP 1 - AIR TEMPERATURE
Air temperature at calibration (C) |  |  | Measure with a thermometer, 0.1 C resolution is plenty
Reference DO at 100% saturation (mg/L) |  |  | Looked up automatically from the Sat Table tab
STEP 2 - MEASURE THE TWO VOLTAGES
Reference point | Known DO | Voltage (V)
Sodium sulfite - zero solution | 0
Probe in air (100% saturation)
Let the probe settle 2 minutes at each point. Read the voltage from the [DO CAL] line in the Serial Monitor.
STEP 3 - RESULT
Slope (mg/L per volt)
Zero offset (V)
PASTE THIS INTO oxyfeeder_firmware.ino
SANITY CHECK
Type any voltage to see what DO it gives
Resulting DO (mg/L)
Put your zero voltage in and you should get 0.00. Put your air voltage in and you should get the reference value.
NOTES
Air calibration works because water-saturated air has the same oxygen partial pressure as water at 100% saturation.
The mg/L value for 100% saturation depends on TEMPERATURE - that is why step 1 exists. There is no temperature-free reference.
At 26-32 C the reference lands between 7.3 and 8.1 mg/L. A value near 13.5 mg/L belongs to water at about 3 C.
This firmware does no temperature compensation, so readings taken far from the calibration temperature carry some error. State that as a limitation.
Mix the sulfite solution fresh each session - once its oxygen is consumed it stops reading zero.

## Sat Table
DO SATURATION TABLE |  | Freshwater, sea level (760 mmHg). Do not edit - other sheets look values up here.
Temp (C) | DO at 100% sat (mg/L) | Source: standard oxygen solubility tables, 0.1 C resolution.
0 | 14.62
0.1 | 14.58
0.2 | 14.54
0.3 | 14.5
0.4 | 14.46
0.5 | 14.42
0.6 | 14.38
0.7 | 14.34
0.8 | 14.3
0.9 | 14.26
1 | 14.22
1.1 | 14.18
1.2 | 14.14
1.3 | 14.1
1.4 | 14.06
1.5 | 14.03
1.6 | 13.99
1.7 | 13.95
1.8 | 13.91
1.9 | 13.87
2 | 13.83
2.1 | 13.79
2.2 | 13.76
2.3 | 13.72
2.4 | 13.68
2.5 | 13.64
2.6 | 13.61
2.7 | 13.57
2.8 | 13.53
2.9 | 13.5
3 | 13.46
3.1 | 13.43
3.2 | 13.39
3.3 | 13.36
3.4 | 13.32
3.5 | 13.29
3.6 | 13.25
3.7 | 13.21
3.8 | 13.18
3.9 | 13.14
4 | 13.11
4.1 | 13.08
4.2 | 13.04
4.3 | 13.01
4.4 | 12.97
4.5 | 12.94
4.6 | 12.91
4.7 | 12.87
4.8 | 12.84
4.9 | 12.8
5 | 12.77
5.1 | 12.74
5.2 | 12.71
5.3 | 12.67
5.4 | 12.64
5.5 | 12.61
5.6 | 12.58
5.7 | 12.55
5.8 | 12.51
5.9 | 12.48
6 | 12.45
6.1 | 12.42
6.2 | 12.39
6.3 | 12.36
6.4 | 12.33
6.5 | 12.29
6.6 | 12.26
6.7 | 12.23
6.8 | 12.2
6.9 | 12.17
7 | 12.14
7.1 | 12.11
7.2 | 12.08
7.3 | 12.05
7.4 | 12.02
7.5 | 11.99
7.6 | 11.96
7.7 | 11.93
7.8 | 11.9
7.9 | 11.87
8 | 11.84
8.1 | 11.81
8.2 | 11.78
8.3 | 11.76
8.4 | 11.73
8.5 | 11.7
8.6 | 11.67
8.7 | 11.64
8.8 | 11.62
8.9 | 11.59
9 | 11.56
9.1 | 11.53
9.2 | 11.51
9.3 | 11.48
9.4 | 11.45
9.5 | 11.43
9.6 | 11.4
9.7 | 11.37
9.8 | 11.34
9.9 | 11.32
10 | 11.29
10.1 | 11.26
10.2 | 11.24
10.3 | 11.21
10.4 | 11.19
10.5 | 11.16
10.6 | 11.13
10.7 | 11.11
10.8 | 11.08
10.9 | 11.06
11 | 11.03
11.1 | 11
11.2 | 10.98
11.3 | 10.95
11.4 | 10.93
11.5 | 10.9
11.6 | 10.88
11.7 | 10.86
11.8 | 10.83
11.9 | 10.8
12 | 10.78
12.1 | 10.76
12.2 | 10.73
12.3 | 10.71
12.4 | 10.68
12.5 | 10.66
12.6 | 10.64
12.7 | 10.61
12.8 | 10.59
12.9 | 10.56
13 | 10.54
13.1 | 10.52
13.2 | 10.49
13.3 | 10.47
13.4 | 10.45
13.5 | 10.43
13.6 | 10.4
13.7 | 10.38
13.8 | 10.36
13.9 | 10.33
14 | 10.31
14.1 | 10.29
14.2 | 10.26
14.3 | 10.24
14.4 | 10.22
14.5 | 10.2
14.6 | 10.17
14.7 | 10.15
14.8 | 10.13
14.9 | 10.1
15 | 10.08
15.1 | 10.06
15.2 | 10.04
15.3 | 10.02
15.4 | 10
15.5 | 9.97
15.6 | 9.95
15.7 | 9.93
15.8 | 9.91
15.9 | 9.89
16 | 9.87
16.1 | 9.85
16.2 | 9.83
16.3 | 9.81
16.4 | 9.79
16.5 | 9.77
16.6 | 9.74
16.7 | 9.72
16.8 | 9.7
16.9 | 9.68
17 | 9.66
17.1 | 9.64
17.2 | 9.62
17.3 | 9.6
17.4 | 9.58
17.5 | 9.57
17.6 | 9.55
17.7 | 9.53
17.8 | 9.51
17.9 | 9.49
18 | 9.47
18.1 | 9.45
18.2 | 9.43
18.3 | 9.41
18.4 | 9.39
18.5 | 9.38
18.6 | 9.36
18.7 | 9.34
18.8 | 9.32
18.9 | 9.3
19 | 9.28
19.1 | 9.26
19.2 | 9.24
19.3 | 9.22
19.4 | 9.2
19.5 | 9.18
19.6 | 9.17
19.7 | 9.15
19.8 | 9.13
19.9 | 9.11
20 | 9.09
20.1 | 9.07
20.2 | 9.06
20.3 | 9.04
20.4 | 9.02
20.5 | 9
20.6 | 8.99
20.7 | 8.97
20.8 | 8.95
20.9 | 8.94
21 | 8.92
21.1 | 8.9
21.2 | 8.88
21.3 | 8.87
21.4 | 8.85
21.5 | 8.83
21.6 | 8.81
21.7 | 8.79
21.8 | 8.78
21.9 | 8.76
22 | 8.74
22.1 | 8.72
22.2 | 8.71
22.3 | 8.69
22.4 | 8.68
22.5 | 8.66
22.6 | 8.64
22.7 | 8.63
22.8 | 8.61
22.9 | 8.6
23 | 8.58
23.1 | 8.56
23.2 | 8.55
23.3 | 8.53
23.4 | 8.52
23.5 | 8.5
23.6 | 8.48
23.7 | 8.47
23.8 | 8.45
23.9 | 8.44
24 | 8.42
24.1 | 8.4
24.2 | 8.39
24.3 | 8.37
24.4 | 8.36
24.5 | 8.34
24.6 | 8.32
24.7 | 8.31
24.8 | 8.29
24.9 | 8.28
25 | 8.26
25.1 | 8.24
25.2 | 8.23
25.3 | 8.21
25.4 | 8.2
25.5 | 8.18
25.6 | 8.17
25.7 | 8.15
25.8 | 8.14
25.9 | 8.12
26 | 8.11
26.1 | 8.1
26.2 | 8.08
26.3 | 8.07
26.4 | 8.05
26.5 | 8.04
26.6 | 8.03
26.7 | 8.01
26.8 | 8
26.9 | 7.98
27 | 7.97
27.1 | 7.96
27.2 | 7.94
27.3 | 7.93
27.4 | 7.91
27.5 | 7.9
27.6 | 7.89
27.7 | 7.87
27.8 | 7.86
27.9 | 7.84
28 | 7.83
28.1 | 7.82
28.2 | 7.8
28.3 | 7.79
28.4 | 7.77
28.5 | 7.76
28.6 | 7.75
28.7 | 7.73
28.8 | 7.72
28.9 | 7.7
29 | 7.69
29.1 | 7.68
29.2 | 7.66
29.3 | 7.65
29.4 | 7.64
29.5 | 7.62
29.6 | 7.61
29.7 | 7.6
29.8 | 7.59
29.9 | 7.57
30 | 7.56
30.1 | 7.55
30.2 | 7.53
30.3 | 7.52
30.4 | 7.51
30.5 | 7.49
30.6 | 7.48
30.7 | 7.47
30.8 | 7.46
30.9 | 7.44
31 | 7.43
31.1 | 7.42
31.2 | 7.41
31.3 | 7.39
31.4 | 7.38
31.5 | 7.37
31.6 | 7.36
31.7 | 7.35
31.8 | 7.33
31.9 | 7.32
32 | 7.31
32.1 | 7.3
32.2 | 7.28
32.3 | 7.27
32.4 | 7.26
32.5 | 7.24
32.6 | 7.23
32.7 | 7.22
32.8 | 7.21
32.9 | 7.19
33 | 7.18
33.1 | 7.17
33.2 | 7.16
33.3 | 7.15
33.4 | 7.14
33.5 | 7.12
33.6 | 7.11
33.7 | 7.1
33.8 | 7.09
33.9 | 7.08
34 | 7.07
34.1 | 7.06
34.2 | 7.05
34.3 | 7.03
34.4 | 7.02
34.5 | 7.01
34.6 | 7
34.7 | 6.99
34.8 | 6.97
34.9 | 6.96
35 | 6.95
35.1 | 6.94
35.2 | 6.93
35.3 | 6.92
35.4 | 6.91
35.5 | 6.89
35.6 | 6.88
35.7 | 6.87
35.8 | 6.86
35.9 | 6.85
36 | 6.84
36.1 | 6.83
36.2 | 6.82
36.3 | 6.81
36.4 | 6.8
36.5 | 6.79
36.6 | 6.77
36.7 | 6.76
36.8 | 6.75
36.9 | 6.74
37 | 6.73
37.1 | 6.72
37.2 | 6.71
37.3 | 6.7
37.4 | 6.69
37.5 | 6.68
37.6 | 6.66
37.7 | 6.65
37.8 | 6.64
37.9 | 6.63
38 | 6.62
38.1 | 6.61
38.2 | 6.6
38.3 | 6.59
38.4 | 6.58
38.5 | 6.57
38.6 | 6.56
38.7 | 6.55
38.8 | 6.54
38.9 | 6.53
39 | 6.52
39.1 | 6.51
39.2 | 6.5
39.3 | 6.49
39.4 | 6.48
39.5 | 6.46
39.6 | 6.45
39.7 | 6.44
39.8 | 6.43
39.9 | 6.42
40 | 6.41